Calculation Formulas

## 1. Inputs and frame

- Our maps: F = front surface, T = thickness, in µm, on a grid of pitch Δ (0.5 or 0.2 mm).

- Pixel centre: x = Δ/2 + Δ·j − 149 mm (same for y); centre (149, 149) from a boundary circle fit. R = √(x² + y²), θ = atan2(y, x).

- Median surface M = F − T/2. Back surface z_b = F − T.

- Evaluation area (FQA): R ≤ 150 − EE − Δ/2, EE = 2 or 3 mm.

- LS(z): least-squares plane c₀ + c₁x + c₂y fitted over the stated area.

- Results in nm are µm × 1000.

## 2. Thickness and global flatness

| Metric | Formula |
| --- | --- |
| THKMean | mean(T) over FQA |
| THKCTR | mean(T) over R < 1 mm |
| GBIR | max(T) − min(T) over FQA |
| GFLR | max(r) − min(r), r = T − LS(T) over FQA |
| GFLD | r at max ¦r¦, sign kept |

## 3. Shape

| Metric | Formula |
| --- | --- |
| Warp-BF | max(m′) − min(m′), m′ = M − LS(M) over FQA |
| Bow-BF | mean(m′) over R < 1 mm |
| Warp-3P | range(M − P₃) over FQA |
| Bow-3P | mean(M − P₃) over R < 1 mm |

P₃ = plane through the pixels nearest R = 150 − EE at θ = 90°, 210°, 330°.

## 4. Site flatness (26 × 8 mm)

- Site box: [x_a, x_a + 26) × [y_a, y_a + 8) mm. Grid placements tested: site corner at the origin, and sites centred on the origin.

- Site selection: “full” = all four corners inside R ≤ 150 − EE; “half” = ≥ 50 % of the box’s pixels in the FQA.

- Within each site s (FQA pixels only):

| Metric | Formula |
| --- | --- |
| SBIR | range(T) |
| SFQR | range(T − LS_s(T)), plane fitted within the site |
| SFLR | range(T − LS_FQA(T)), global plane |

Reported as maximum and mean over sites.

## 5. Edge sectors

- 72 sectors × 5°, starting at 0°; radial band R ∈ [R_f − L, R_f], R_f = 150 − EE, L = 30 or 35 mm.

| Metric | Formula |
| --- | --- |
| ESBIR | range(T) in each sector |
| ESFQR | range(T − LS_sector(T)) |

Reported as maximum and mean over sectors.

## 6. Diagnostic tests

- Clipped pixels: pixels where T equals exactly the map’s global maximum or minimum.

- Texture smoothing: T_σ = (g_σ ∗ (T·v)) / (g_σ ∗ v), g_σ Gaussian with σ = 0.5, 1, 2 mm, v = valid-pixel mask. SFQR, SBIR, ESFQR recomputed on T_σ.

- Shape decomposition: fit M = c₀ + c₁x + c₂y + c₃·r̂² + c₄·(x² − y²)/148² + c₅·2xy/148², r̂ = R/148. c₃ = bowl term; c₄, c₅ = saddle terms.

- Missing-bowl test: M_A = M + A·(R/148)², A raised until Bow-BF matches WS2+ (A = 2.25 µm); Warp-BF, Bow-3P, Warp-3P then recomputed.

- Texture spectrum: z = T minus a 6-term quadratic fit within R ≤ 147 mm; W = 2-D Hann × mask; P(k) = |FFT₂(z·W)|² · Δ² / ΣW². Band rms = √(Σ_band P · Δk²), Δk = 1/(N·Δ). Directionality = share of power at k > 0.3 cycles/mm within ±10° of the strongest direction.

## 7. Nanotopography — SQMM emulation (0.2 mm map)

- Edge-safe Gaussian: G_σ(z) = (g_σ ∗ (z·v)) / (g_σ ∗ v), kernel truncated at 3σ.

- Double-Gaussian high-pass: HP_λ(z) = (1 − G_σ)² z, with σ = λ·√(1.228 / 2π²) ≈ 0.2495·λ, which gives 50 % transmission at wavelength λ: (1 − e^(−2π²σ²/λ²))² = 0.5. This form cancels wafer curvature exactly.

- Shrinking filter: d = distance to the wafer edge (Euclidean distance transform). Local cutoff λ(d) = clip(d / (3·√2·0.2495), 1 mm, 20 mm), keeping ±3σ of the combined kernel on the wafer. z_NT computed at 11 cutoffs spaced geometrically from 1 to 20 mm and linearly interpolated at λ(d).

- Window PV: square window of side L = 0.5, 2, 10 mm (k = round(L/Δ) pixels, forced odd); PV(x, y) = max − min of z_NT within the window. Only windows entirely inside R ≤ 147 mm (EE = 3 mm) are kept (partial windows off).

- THA: THA_P = (100 − P)th percentile of PV over all window centres; P0.05 → 99.95th percentile. P0 = max(PV). Partial-window test: repeat with every window centre in the FQA.

- Clean-up variants:

  - Despike: if |z − median₅ₓ₅(z)| > 6 nm, replace the pixel with the median.

  - De-grid: z_NT − c(x) − r(y), where c and r are the column and row means of z_NT over the FQA, each minus its own 15-pixel (3 mm) moving average.

- 0.05 % vs 5 % test: compare P0.05/P0 computed with the 99.95th percentile and with the 95th percentile against the WS2+ ratios.

## 8. Fringe height step

- 12.8 mm patches of the surface h (front F or back z_b) and of its z_NT.

- slope = |mean ∇h| (nm/mm).

- Fringe period = 1/|k_peak|, from the 2-D FFT of the Hann-windowed z_NT patch, ignoring |k| < 0.15 cycles/mm.

- Height step per fringe = slope × period.

- A patch is kept if slope > 15 nm/mm and |cos(k_peak, ∇h)| > 0.8 (fringes aligned with the slope).

## 9. Assumptions

- Exact shrink law of the WS2+ shrinking filter (here: kernel ±3σ kept on the wafer, cutoff 20 → 1 mm).

- SQMM internal cutoff and kernel (here: 20 mm; 10 mm gives the same THA within 3 %).

- WS2+ 3-point reference positions for Bow-3P / Warp-3P (here: 90°, 210°, 330° at R = 150 − EE).

- WS2+ site-grid origin and partial-site rule (both placements reported).