# Calculation Formulas — v2

WS2+-class geometry analysis · engine 0.7.0 · project document v1.6 §11 / §13 · 2026-10-13

Changes from v1 are listed in section 10. Results in nm are µm × 1000. Frequencies in cycles/mm.

## 1. Inputs and frame

- Maps: F = front surface, T = thickness, in µm, on a grid of pitch Δ (0.5 mm global geometry, 0.2 mm nanotopography).

- Centre: circle fit to the outer boundary of the measured disc (holes filled first; radius corrected by Δ/2 because boundary pixel centres lie half a pixel inside the edge). Pixel centre x = Δ/2 + Δ·j − x_c (same for y). For the project-team export the fitted centre is (149, 149) mm in the file frame. R = √(x² + y²), θ = atan2(y, x).

- Sign: +z out of the front surface. Median surface M = F − T/2 = (z_front + z_back)/2. Back surface z_b = F − T.

- Evaluation area (FQA): pixel centres with R ≤ 150 − EE. EE is taken from the recipe (2 or 3 mm; SQMM fixed at 3 mm).

- LS(z): least-squares plane c₀ + c₁x + c₂y fitted over the stated area.

- Centre value C(z): mean of z over valid pixels with R < 1 mm (pixels exactly on R = 1 mm excluded). If fewer than half of the pixels inside 1 mm are valid, C is not reported. On grids with no pixel inside 1 mm, C is the mean of the nearest valid pixels.

## 2. Thickness and global flatness

| Metric | Formula |
| --- | --- |
| THKMean | mean(T) over the FQA |
| THKCTR | C(T) |
| GBIR | max(T) − min(T) over the FQA |
| GFLR | max(r) − min(r), r = T − LS(T) over the FQA |
| GFLD | r at max ¦r¦, sign kept |

## 3. Shape

| Metric | Formula |
| --- | --- |
| Warp-BF | max(m′) − min(m′), m′ = M − LS(M) over the FQA |
| Bow-BF | C(m′) |
| Warp-3P | range(M − P₃) over the FQA |
| Bow-3P | C(M − P₃) |

P₃ = plane through the valid pixels nearest R = 150 − EE at θ = 90°, 210°, 330°.

Sign: Bow is positive when the centre of the median surface lies above the reference plane (front side up).

## 4. Site flatness (26 × 8 mm)

- Site box: [x_a, x_a + 26) × [y_a, y_a + 8) mm. Grid origin is a recipe parameter (site corner at the origin, or sites centred on the origin) until the WS2+ origin is confirmed.

- Site selection: “full” = all four corners inside R ≤ 150 − EE; “partial” = FQA pixels ≥ 50 % of the full box area (Δ-pixel count of the 26 × 8 mm box, including any part outside the map).

- Within each site s (FQA pixels only):

| Metric | Formula |
| --- | --- |
| SBIR | range(T) |
| SFQR | range(T − LS_s(T)), plane fitted within the site |
| SFLR | range(T − LS_FQA(T)), global plane |

Reported as maximum and mean over sites, with the per-site list.

## 5. Edge sectors

- 72 sectors × 5°, starting at 0°; radial band R ∈ [R_f − L, R_f], R_f = 150 − EE, L = 30 or 35 mm. Sectors with fewer than 3 pixels are not evaluated.

| Metric | Formula |
| --- | --- |
| ESBIR | range(T) in each sector |
| ESFQR | range(T − LS_sector(T)) |

Reported as maximum and mean over sectors, with the per-sector list.

## 6. Diagnostic tests (analysis only — not reported metrics)

- Clipped pixels: pixels pinned at the map’s global maximum or minimum that lie more than 3 pixels inside the measured disc. (A smooth tilted map reaches its extreme on the boundary; a plateau inside the disc means the export clipped the data.) Range metrics of a clipped map are lower bounds.

- Texture smoothing: T_σ = (g_σ ∗ (T·v)) / (g_σ ∗ v), g_σ Gaussian with σ = 0.5, 1, 2 mm, v = valid-pixel mask. SFQR, SBIR, ESFQR recomputed on T_σ.

- Shape decomposition: fit M = c₀ + c₁x + c₂y + c₃·r̂² + c₄·(x² − y²)/148² + c₅·2xy/148², r̂ = R/148. c₃ = bowl term; c₄, c₅ = saddle terms.

- Missing-bowl test: M_A = M + A·(R/148)², A raised until Bow-BF matches WS2+ (A = 2.25 µm on the reference wafer); Warp-BF, Bow-3P, Warp-3P then recomputed.

- Texture spectrum: z = T minus a 6-term quadratic fit within R ≤ 147 mm; W = 2-D Hann × mask; P(k) = |FFT₂(z·W)|² · Δ² / ΣW². Band rms = √(Σ_band P · Δk²), Δk = 1/(N·Δ). Directionality = share of power at k > 0.3 cycles/mm within ±10° of the strongest direction.

## 7. Nanotopography — Gaussian and SQMM recipes

- Edge-safe Gaussian: G_σ(z) = (g_σ ∗ (z·v)) / (g_σ ∗ v), kernel truncated at 3σ.

- Double-Gaussian high-pass: HP_λ(z) = (1 − G_σ)² z, with σ = λ·√(1.228 / 2π²) ≈ 0.2495·λ, which gives 50 % transmission at wavelength λ: (1 − e^(−2π²σ²/λ²))² = 0.5. In the interior (kernel entirely on the wafer) it cancels planes and quadratic curvature exactly; near the edge the normalised Gaussian is only approximately so. Transmission values quoted are for the ideal (untruncated) Gaussian; the 3σ truncation passes about 2–4 % less at 1.5–3 × the cutoff wavelength.

- Shrinking filter: d = distance to the outer edge of the evaluation area, after filling holes such as pixel dropouts (Euclidean distance transform). Local cutoff λ(d) = clip(d / (3·√2·0.2495), λ_edge, λ_centre), keeping ±3σ of the combined kernel (σ√2) on the wafer; default 20 → 1 mm. z_NT computed at 11 cutoffs (recipe filter.levels) spaced geometrically from λ_edge to λ_centre and interpolated linearly at λ(d). Constant filter when λ_edge = λ_centre.

- Window PV: window of size L: square (side L; SQMM 0.5, 2, 10 mm) or circle (diameter L; Gaussian recipes, e.g. 2 and 10 mm). k = round(L/Δ) pixels, forced odd (a 2 mm window at 0.2 mm is 11 pixels, spanning 2.0 mm centre to centre). Windows smaller than 2 pixels are not evaluated. PV(x, y) = max − min of z_NT within the window.

- Partial windows: off — only windows whose whole footprint lies inside the FQA (R ≤ 147 mm for EE 3 mm) are kept.

- THA: THA_P = (100 − P)th percentile of PV over all kept window centres; P = 0.05 % of area → 99.95th percentile. P0 = max(PV). Confirmed: P0.05 means 0.05 % of area (not 5 %), from the WS2+ P0.05/P0 ratios; partial windows change THA by ≤ 1 % front and ≤ 7 % back.

- SQA: percent of window centres with PV above the recipe’s Max Failed threshold. Avalanche curve: % of area with PV above each threshold.

- Diagnostic clean-up (analysis only, not part of THA): despike — if |z − median₅ₓ₅(z)| > 6 nm, replace the pixel with the median; de-grid — z_NT − c(x) − r(y), c and r the column and row means of z_NT over the FQA, each minus its own 15-pixel (3 mm) moving average.

## 8. Fringe height step (diagnostic)

- 12.8 mm patches of the surface h (front F or back z_b) and of its z_NT.

- slope = |mean ∇h| (nm/mm).

- Fringe period = 1/|k_peak|, from the 2-D FFT of the Hann-windowed z_NT patch, ignoring |k| < 0.15 cycles/mm.

- Height step per fringe = slope × period.

- A patch is kept if slope > 15 nm/mm and |cos(k_peak, ∇h)| > 0.8 (fringes aligned with the slope).

## 9. Open assumptions (recipe parameters with defaults)

- Shrink law of the WS2+ shrinking filter, and whether d is measured to the wafer edge or to the EE boundary (here: EE boundary, kernel ±3σ kept on the wafer, 20 → 1 mm, 11 levels).

- SQMM internal cutoff and kernel (here: 20 mm; 10 mm gives the same THA within 3 %).

- WS2+ 3-point reference positions for Bow-3P / Warp-3P (here: 90°, 210°, 330° at R = 150 − EE).

- WS2+ site-grid origin and partial-site rule.

- ZDD: not implemented (SEMI M68 reference fit not fixed in v1.6).

- PSD band boundary k_low: 0.02 cycles/mm (50 mm, SEMI) or 0.05 cycles/mm (20 mm); bands are read from configuration.

## 10. Changes from v1

| # | Section | Change |
| --- | --- | --- |
| 1 | 1 | Centre by circle fit (holes filled, Δ/2 radius correction) instead of a fixed (149, 149); FQA = R ≤ 150 − EE on pixel centres (v1 had an extra −Δ/2 margin). |
| 2 | 1–3 | Centre value C(z) defined once: mean over R < 1 mm, boundary excluded, not reported if under half valid. The engine used the single nearest pixel until 0.6; changed in engine 0.7.0 (metric change: THK centre −2.7 to −4.2 nm, Bow +2.8 to +3.9 nm on the reference wafer). |
| 3 | 3 | Bow sign convention stated. |
| 4 | 4 | Partial-site rule uses the full box area; “half” renamed “partial”. |
| 5 | 6 | Clipping counted only as plateaus inside the disc; section marked diagnostic. |
| 6 | 7 | “Cancels curvature exactly” limited to the interior; ideal vs truncated transmission noted. |
| 7 | 7 | Shrinking-filter distance to the outer FQA edge with holes filled; 11 levels as a recipe setting. |
| 8 | 7 | Circle windows added; window-size rounding and the 2-pixel minimum stated. |
| 9 | 7 | Despike and de-grid marked as diagnostics; the 0.05 % result recorded; SQA and avalanche added. |
| 10 | 9 | k_low, ZDD and the shrink-distance question added to the open assumptions. |